VALIDATION REPORT
OPERATIONAL QUALIFICATION
MSA MODULE

JR Validated Environment — MSA Module v1.0

| Field | Value |
| --- | --- |
| Document Number | JR-VR-MSA-001 |
| Title | Validation Report — Operational Qualification, MSA Module |
| Validation Plan | JR-VP-MSA-001 v1.0 |
| System | JR Validated Environment — MSA Module |
| Module Version | 1.0 |
| Document Version | 1.0 |
| Execution Date | 2026-03-18 |
| OQ Result | PASS — 53 / 53 tests passed, 0 failures, 0 deviations |
| Author | Joep Rous |
| Reviewer | [Name] |
| Approver | [Name] |

# Version History

| Version | Date | Author | Description |
| --- | --- | --- | --- |
| 1.0 | 2026-03-18 | Joep Rous | Initial OQ execution report. All 53 test cases passed. No deviations. |

# Approval Signatures

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| Author | Joep Rous |  | 2026-03-18 |
| Reviewer | [Name] |  |  |
| Approver (QA) | [Name] |  |  |

# 1.  Purpose

This document reports the execution and outcome of the Operational Qualification (OQ) test suite for the JR Validated Environment MSA Module, version 1.0. It provides objective evidence that all MSA scripts perform as specified in the OQ Validation Plan JR-VP-MSA-001 v1.0 under the defined test conditions.

This report records the test environment, execution results for all 53 test cases, and the formal OQ conclusion.

# 2.  Scope

## 2.1  In Scope

All 5 R scripts in repos/msa/R/ (jrc_msa_gauge_rr, jrc_msa_nested_grr, jrc_msa_linearity_bias, jrc_msa_type1, jrc_msa_attribute)

Correct computation of outputs for valid inputs

Correct rejection of invalid inputs with informative error messages

Bypass-protection verification (RENV_PATHS_ROOT check)

PNG output to ~/Downloads/

## 2.2  Out of Scope

Demo scripts jrc_msa_R_hello.R and jrc_msa_py_hello.py

Infrastructure scripts in bin/ and admin/ (covered by JR-IQ-001)

Performance qualification (PQ) with real production data

Validation of R interpreter and third-party packages (ggplot2, grid)

# 3.  Reference Documents

| Document ID | Title | Location |
| --- | --- | --- |
| JR-VP-MSA-001 v1.0 | Validation Plan — OQ, MSA Module | repos/msa/docs/msa_validation_plan.docx |
| JR-VP-002 v1.0 | Validation Plan — OQ, Community Script Suite v1.1.0 | docs/ignore/oq_validation_plan.docx |
| JR-IQ-001 | IQ Execution Evidence | docs/IQ_validation_20260311_205146.txt |
| OQ Evidence | MSA OQ Execution Evidence (pytest output + header) | ~/.jrscript/MyProject/validation/msa_oq_execution_20260318T150625.txt |

# 4.  Test Environment

| Item | Value |
| --- | --- |
| Execution date / time | 2026-03-18T15:06:25 |
| Completion date / time | 2026-03-18T15:06:47 |
| Host | Joeps-MacBook-Air.local |
| Operating system | macOS 15.7.3 (Darwin 24.6.0, aarch64) |
| R version | 4.5.0 |
| Python version | 3.11.9 |
| pytest version | 8.3.4 |
| OQ virtual environment | ~/.venvs/MyProject_oq/ (reused from core OQ suite) |
| Script entry point | bin/jrrun (sets RENV_PATHS_ROOT, loads renv library) |
| Test runner | repos/msa/admin_msa_oq |
| Test directory | repos/msa/oq/ |
| Test data directory | repos/msa/oq/data/ (22 committed synthetic CSV files) |
| Evidence file | ~/.jrscript/MyProject/validation/msa_oq_execution_20260318T150625.txt |

# 5.  Test Data

All test data files are synthetic CSV files committed to repos/msa/oq/data/ and included in the project integrity check (admin/project_integrity.sha256). No real patient or production data is used.

| File | Content | Used by |
| --- | --- | --- |
| gauge_rr_balanced.csv | 10 parts × 3 operators × 3 reps, low noise | TC-MSA-GRR-001..004 |
| gauge_rr_missing_col.csv | Missing 'operator' column | TC-MSA-GRR-007 |
| gauge_rr_unbalanced.csv | Unequal replicates per cell | TC-MSA-GRR-008 |
| gauge_rr_one_operator.csv | Single operator only | TC-MSA-GRR-009 |
| nested_grr_good.csv | 5 parts × 3 operators × 2 reps, low noise, nested design | TC-MSA-NGR-001..005 |
| nested_grr_poor.csv | High within-operator replicate variation (large EV) | TC-MSA-NGR-003 |
| nested_grr_missing_col.csv | Missing 'replicate' column | TC-MSA-NGR-008 |
| nested_grr_one_operator.csv | Single operator only | TC-MSA-NGR-009 |
| nested_grr_unbalanced.csv | Operator A has 3 parts, operator B has 2 parts | TC-MSA-NGR-010 |
| linearity_bias_good.csv | 5 reference levels × 3 reps, slope ≈ 0.05 | TC-MSA-LB-001..004 |
| linearity_bias_missing_col.csv | Missing 'reference' column | TC-MSA-LB-007 |
| linearity_bias_one_part.csv | Single reference level (cannot fit regression) | TC-MSA-LB-008 |
| linearity_bias_inconsistent_ref.csv | Same part ID mapped to two different reference values | TC-MSA-LB-009 |
| type1_good.csv | 25 measurements near reference (50.0), low bias | TC-MSA-T1-001..004 |
| type1_biased.csv | 25 measurements with systematic positive bias ≈ 0.15 | TC-MSA-T1-003 |
| type1_missing_value_col.csv | Missing 'value' column | TC-MSA-T1-009 |
| type1_too_few.csv | Only 5 measurements (< 10 required) | TC-MSA-T1-010 |
| attribute_with_ref.csv | 3 appraisers × 10 parts × 2 trials, with reference column; appraiser A is perfect vs reference | TC-MSA-ATT-001, 003..005 |
| attribute_no_ref.csv | Same design without reference column | TC-MSA-ATT-002 |
| attribute_missing_col.csv | Missing 'trial' column | TC-MSA-ATT-008 |
| attribute_one_appraiser.csv | Single appraiser only | TC-MSA-ATT-009 |
| attribute_unbalanced.csv | Unequal trials per appraiser-part cell | TC-MSA-ATT-010 |

# 6.  Test Execution Results

## 6.1  Summary

| Item | Value |
| --- | --- |
| Total test cases | 53 |
| Passed | 53 |
| Failed | 0 |
| Errors | 0 |
| Duration | 22.45 s |
| OQ result | PASS |

## 6.2  Results by Test File

| Test file | Script(s) covered | Tests | Result |
| --- | --- | --- | --- |
| test_msa_attribute.py | jrc_msa_attribute | 11 | 11 / 11 PASS |
| test_msa_gauge_rr.py | jrc_msa_gauge_rr | 10 | 10 / 10 PASS |
| test_msa_linearity_bias.py | jrc_msa_linearity_bias | 10 | 10 / 10 PASS |
| test_msa_nested_grr.py | jrc_msa_nested_grr | 11 | 11 / 11 PASS |
| test_msa_type1.py | jrc_msa_type1 | 11 | 11 / 11 PASS |

## 6.3  Individual Test Case Results

All 53 test cases passed. The full pytest output, including individual test IDs, pass/fail status, and timing, is recorded in the evidence file:

~/.jrscript/MyProject/validation/msa_oq_execution_20260318T150625.txt

That file also contains the evidence header (host, R version, Python version, pytest version, timestamp) and the OQ RESULT: PASS footer. The per-test results are reproduced in the table below for traceability.

| Test Case ID | Description | Expected | Result |
| --- | --- | --- | --- |
| TC-MSA-GRR-001 | Happy path — exit 0, sections present | Exit 0, sections present | PASS |
| TC-MSA-GRR-002 | Known data — %GRR < 10%, verdict ACCEPTABLE | %GRR < 10%, ACCEPTABLE | PASS |
| TC-MSA-GRR-003 | --tolerance flag | Exit 0, tolerance output | PASS |
| TC-MSA-GRR-004 | PNG written to ~/Downloads/ | PNG present, recent mtime | PASS |
| TC-MSA-GRR-005 | No arguments → usage message | Exit ≠ 0, 'Usage' in output | PASS |
| TC-MSA-GRR-006 | File not found | Exit ≠ 0, 'not found' in output | PASS |
| TC-MSA-GRR-007 | Missing column → column named | Exit ≠ 0, column name in output | PASS |
| TC-MSA-GRR-008 | Unbalanced design | Exit ≠ 0, 'unbalanced' in output | PASS |
| TC-MSA-GRR-009 | Single operator | Exit ≠ 0 | PASS |
| TC-MSA-GRR-010 | Bypass protection — direct Rscript call | Exit ≠ 0, 'RENV_PATHS_ROOT' | PASS |
| TC-MSA-NGR-001 | Happy path — exit 0, sections present | Exit 0, sections present | PASS |
| TC-MSA-NGR-002 | Known good data — %GRR < 30%, ACCEPTABLE/MARGINAL | %GRR < 30% | PASS |
| TC-MSA-NGR-003 | Known poor data — verdict UNACCEPTABLE | UNACCEPTABLE in output | PASS |
| TC-MSA-NGR-004 | --tolerance flag | Exit 0, tolerance output | PASS |
| TC-MSA-NGR-005 | PNG written to ~/Downloads/ | PNG present, recent mtime | PASS |
| TC-MSA-NGR-006 | No arguments → usage message | Exit ≠ 0, 'Usage' in output | PASS |
| TC-MSA-NGR-007 | File not found | Exit ≠ 0, 'not found' in output | PASS |
| TC-MSA-NGR-008 | Missing 'replicate' column | Exit ≠ 0, 'replicate' in output | PASS |
| TC-MSA-NGR-009 | Single operator | Exit ≠ 0 | PASS |
| TC-MSA-NGR-010 | Unbalanced design | Exit ≠ 0, 'unbalanced' in output | PASS |
| TC-MSA-NGR-011 | Bypass protection — direct Rscript call | Exit ≠ 0, 'RENV_PATHS_ROOT' | PASS |
| TC-MSA-LB-001 | Happy path — exit 0, sections present | Exit 0, sections present | PASS |
| TC-MSA-LB-002 | Known data — slope significant | Slope in [0.01, 0.15] | PASS |
| TC-MSA-LB-003 | --tolerance flag | Exit 0, tolerance output | PASS |
| TC-MSA-LB-004 | PNG written to ~/Downloads/ | PNG present, recent mtime | PASS |
| TC-MSA-LB-005 | No arguments → usage message | Exit ≠ 0, 'Usage' in output | PASS |
| TC-MSA-LB-006 | File not found | Exit ≠ 0, 'not found' in output | PASS |
| TC-MSA-LB-007 | Missing 'reference' column | Exit ≠ 0, column named | PASS |
| TC-MSA-LB-008 | Single reference level | Exit ≠ 0 | PASS |
| TC-MSA-LB-009 | Inconsistent reference per part | Exit ≠ 0 | PASS |
| TC-MSA-LB-010 | Bypass protection — direct Rscript call | Exit ≠ 0, 'RENV_PATHS_ROOT' | PASS |
| TC-MSA-T1-001 | Happy path — exit 0, Cg/Cgk present | Exit 0, sections present | PASS |
| TC-MSA-T1-002 | Known good data — Cg ≥ 1.33, CAPABLE | Cg ≥ 1.33, Cgk ≥ 1.33, CAPABLE | PASS |
| TC-MSA-T1-003 | Biased data — Cgk < Cg | Cgk < Cg | PASS |
| TC-MSA-T1-004 | PNG written to ~/Downloads/ | PNG present, recent mtime | PASS |
| TC-MSA-T1-005 | No arguments → usage message | Exit ≠ 0, 'Usage' in output | PASS |
| TC-MSA-T1-006 | --reference missing | Exit ≠ 0, 'reference' in output | PASS |
| TC-MSA-T1-007 | --tolerance missing | Exit ≠ 0, 'tolerance' in output | PASS |
| TC-MSA-T1-008 | File not found | Exit ≠ 0, 'not found' in output | PASS |
| TC-MSA-T1-009 | Missing 'value' column | Exit ≠ 0, 'value' in output | PASS |
| TC-MSA-T1-010 | Too few measurements (< 10) | Exit ≠ 0 | PASS |
| TC-MSA-T1-011 | Bypass protection — direct Rscript call | Exit ≠ 0, 'RENV_PATHS_ROOT' | PASS |
| TC-MSA-ATT-001 | Dataset with reference — exit 0, all sections | Exit 0, all sections present | PASS |
| TC-MSA-ATT-002 | Dataset without reference — no vs-ref section | Exit 0, 'Vs Reference' absent | PASS |
| TC-MSA-ATT-003 | Known data — Fleiss Kappa in [0.7, 1.0], verdict | Kappa ∈ [0.7,1.0], verdict | PASS |
| TC-MSA-ATT-004 | Perfect appraiser A vs reference — Kappa = 1.0 | Kappa = 1.0000 for appraiser A | PASS |
| TC-MSA-ATT-005 | PNG written to ~/Downloads/ | PNG present, recent mtime | PASS |
| TC-MSA-ATT-006 | No arguments → usage message | Exit ≠ 0, 'Usage' in output | PASS |
| TC-MSA-ATT-007 | File not found | Exit ≠ 0, 'not found' in output | PASS |
| TC-MSA-ATT-008 | Missing 'trial' column | Exit ≠ 0, 'trial' in output | PASS |
| TC-MSA-ATT-009 | Only one appraiser | Exit ≠ 0 | PASS |
| TC-MSA-ATT-010 | Unbalanced design | Exit ≠ 0, 'unbalanced' in output | PASS |
| TC-MSA-ATT-011 | Bypass protection — direct Rscript call | Exit ≠ 0, 'RENV_PATHS_ROOT' | PASS |

# 7.  Deviations from OQ Plan (JR-VP-MSA-001 v1.0)

No deviations from the approved OQ plan were identified during this OQ execution.

| ID | Description | Resolution | Status |
| --- | --- | --- | --- |
| — | None | — | N/A |

# 8.  Requirements Traceability Matrix

The table below maps each User Requirement from JR-VP-MSA-001 to the test cases executed and their result.

| UR | Requirement (summary) | Test Cases | Result |
| --- | --- | --- | --- |
| UR-MSA-001 | jrc_msa_gauge_rr — ANOVA, variance components, %GRR | TC-MSA-GRR-001..003 | 3 / 3 PASS |
| UR-MSA-002 | Gauge R&R classification (ACCEPTABLE/MARGINAL/UNACCEPTABLE) | TC-MSA-GRR-002 | 1 / 1 PASS |
| UR-MSA-003 | Gauge R&R four-panel PNG to ~/Downloads/ | TC-MSA-GRR-004 | 1 / 1 PASS |
| UR-MSA-004 | Gauge R&R — reject invalid inputs | TC-MSA-GRR-007..009 | 3 / 3 PASS |
| UR-MSA-005 | jrc_msa_nested_grr — nested ANOVA, %GRR | TC-MSA-NGR-001, 002, 004 | 3 / 3 PASS |
| UR-MSA-006 | Nested GRR classification | TC-MSA-NGR-002..003 | 2 / 2 PASS |
| UR-MSA-007 | Nested GRR two-panel PNG to ~/Downloads/ | TC-MSA-NGR-005 | 1 / 1 PASS |
| UR-MSA-008 | Nested GRR — reject invalid inputs | TC-MSA-NGR-008..010 | 3 / 3 PASS |
| UR-MSA-009 | jrc_msa_linearity_bias — regression, slope, %Linearity | TC-MSA-LB-001..003 | 3 / 3 PASS |
| UR-MSA-010 | Linearity/bias verdict | TC-MSA-LB-002, 004 | 2 / 2 PASS |
| UR-MSA-011 | Linearity/bias two-panel PNG to ~/Downloads/ | TC-MSA-LB-004 | 1 / 1 PASS |
| UR-MSA-012 | Linearity/bias — reject invalid inputs | TC-MSA-LB-007..009 | 3 / 3 PASS |
| UR-MSA-013 | jrc_msa_type1 — Cg, Cgk, bias t-test | TC-MSA-T1-001..003 | 3 / 3 PASS |
| UR-MSA-014 | Type 1 verdict; Cgk < Cg when bias present | TC-MSA-T1-002..003 | 2 / 2 PASS |
| UR-MSA-015 | Type 1 two-panel PNG to ~/Downloads/ | TC-MSA-T1-004 | 1 / 1 PASS |
| UR-MSA-016 | Type 1 — mandatory args; reject < 10 measurements | TC-MSA-T1-006..010 | 5 / 5 PASS |
| UR-MSA-017 | jrc_msa_attribute — within/between Kappa | TC-MSA-ATT-001..003 | 3 / 3 PASS |
| UR-MSA-018 | Attribute — vs-reference Kappa when column present | TC-MSA-ATT-001, 004 | 2 / 2 PASS |
| UR-MSA-019 | Attribute Kappa classification and verdict | TC-MSA-ATT-003 | 1 / 1 PASS |
| UR-MSA-020 | Attribute two-panel PNG to ~/Downloads/ | TC-MSA-ATT-005 | 1 / 1 PASS |
| UR-MSA-021 | Attribute — reject invalid inputs | TC-MSA-ATT-008..010 | 3 / 3 PASS |
| UR-MSA-022 | All scripts — RENV_PATHS_ROOT bypass protection | TC-MSA-GRR-010, NGR-011, LB-010, T1-011, ATT-011 | 5 / 5 PASS |
| UR-MSA-023 | All scripts — no-arguments usage message | TC-MSA-GRR-005, NGR-006, LB-005, T1-005, ATT-006 | 5 / 5 PASS |
| UR-MSA-024 | All scripts — file not found error | TC-MSA-GRR-006, NGR-007, LB-006, T1-008, ATT-007 | 5 / 5 PASS |

# 9.  OQ Conclusion

The Operational Qualification of the JR Validated Environment MSA Module v1.0 is complete. All 53 test cases defined in JR-VP-MSA-001 were executed on 2026-03-18 and passed with no failures, no errors, and no deviations.

All 24 user requirements (UR-MSA-001 through UR-MSA-024) are covered by at least one passing test case. The OQ acceptance criteria stated in JR-VP-MSA-001 Section 10 are fully met.

The MSA module is hereby declared OPERATIONALLY QUALIFIED and released for use in design verification activities in accordance with the JR Validated Environment quality system.

OQ Evidence file: ~/.jrscript/MyProject/validation/msa_oq_execution_20260318T150625.txt