JR VALIDATED ENVIRONMENT

Validation Plan

v1.0

| Document ID | VP-JR-001 |
| --- | --- |
| Version | 1.0 |
| Status | [ ] Draft [ ] Under Review [ X] Approved |
| Date | 2026-03-12 |
| Organisation | Dwylup |
| Project | JR Validated Environment |
| Software Version | v1.0 |
| Prepared by | Joep Rous Signature: _______________ Date:2026-03-12 |
| Reviewed by | Joep Rous Signature: _______________ Date: 2026-03-12 |
| Approved by | Joep Rous Signature: _______________ Date: 2026-03-12 |

Regulatory References: FDA 21 CFR Part 11  |  ISO 13485  |  GAMP 5  |  IEC 62304  |  EU MDR 2017/745

# 1. Purpose and Scope

This Validation Plan defines the approach, responsibilities, and test protocol for validating the JR Validated Environment v1.0. The JR Validated Environment provides a controlled, reproducible execution environment for R and Python scripts used in medical device development and regulatory submission workflows.

This document covers Installation Qualification (IQ), Operational Qualification (OQ), and Performance Qualification (PQ) for the core environment. Community scripts are outside the scope of this document.

## 1.1 Regulatory Basis

This validation plan has been prepared in accordance with:

- FDA 21 CFR Part 11 — Electronic records and signatures

- ISO 13485:2016 — Quality management systems for medical devices

- GAMP 5 — A Risk-Based Approach to Compliant GxP Computerised Systems

- IEC 62304:2006+AMD1:2015 — Medical device software lifecycle processes

- EU MDR 2017/745 — EU Medical Device Regulation

## 1.2 Software Classification

Under GAMP 5, the JR Validated Environment is classified as Category 4 (Configured Software). Under IEC 62304, scripts executing within the environment are classified according to the safety class of the device they support. The environment itself is infrastructure software.

## 1.3 Scope Exclusions

- data/standard/ shared dataset library (deferred to v1.1)

- Community scripts and contrib packages

- Windows and Linux platform support (macOS only for v1.0)

# 2. System Description

The JR Validated Environment consists of:

- A controlled local package repository (R and Python) distributed via Dropbox

- Isolated per-project R library and Python virtual environment stored in the user home directory

- A unified wrapper script (jrrun) that enforces environment isolation, verifies project integrity, and auto-rebuilds the environment when package versions change

- Administrative scripts for installation, validation, and environment management

- A project integrity SHA-256 hash file covering all controlled scripts

## 2.1 Key Components

| Component | Location | Purpose |
| --- | --- | --- |
| jrrun | bin/ | Unified wrapper — enforces isolation, integrity check, auto-rebuild |
| admin_install_R | admin/ | Installs R packages from local repo into isolated library |
| admin_install_Python | admin/ | Installs Python packages from local repo into virtual environment |
| admin_validate | admin/ | Generates and runs IQ validation scripts, produces evidence file |
| jr_animate / jr_static | wrapper/ | User entry points — thin shims calling jrrun for hello-world scripts |
| R_repo / Python_repo | Dropbox (external) | Controlled local package repositories — sole package source |

## 2.2 Environment Isolation Mechanism

R packages are installed into ~/.renv/[PROJECT_ID]/library/. Python packages are installed into ~/.venvs/[PROJECT_ID]/. Both are isolated from the system R and Python installations. Direct invocation of scripts outside of jrrun is blocked by explicit environment checks at script startup.

## 2.3 Known Limitations and Asymmetries

| Item | Description | Risk / Mitigation |
| --- | --- | --- |
| R version pinning | miniCRAN --rebuild and --add fetch current CRAN versions, not pinned versions. R_requirements.txt reflects what is in the repo. | Low. The local repository is the controlled artefact. Rebuilding requires re-validation. Documented in admin manual. |
| Python bypass message | Python scripts fail at import time outside jrrun rather than with an explicit error referencing the wrapper. | Low. Scripts cannot run without the validated packages. Improvement deferred to v1.1. |

# 3. Roles and Responsibilities

| Role | Responsibilities | Name / Sign-off |
| --- | --- | --- |
| Validation Lead | Prepares and executes validation plan, records results, authors validation report | [Name] |
| System Administrator | Installs and configures the environment, runs admin_validate, maintains package repository | [Name] |
| Quality Manager | Reviews and approves validation plan and report | [Name] |

# 4. Installation Qualification (IQ)

The IQ confirms that the JR Validated Environment has been correctly installed on the target machine and that all components are present at the correct versions. The IQ is automated via the admin_validate script.

## 4.1 IQ Prerequisites

- macOS with Apple Silicon (arm64) processor

- R 4.5.x installed

- Python 3.11.x installed

- Dropbox synced with R_repo/ and Python_repo/ present

- admin_create_hash has been run to generate project_integrity.sha256

## 4.2 IQ Procedure

| Step | Action | Expected Result | Actual Result |
| --- | --- | --- | --- |
| 1 | Run ./admin/admin_install_R from project root | All R packages installed at pinned versions in ~/.renv/[PROJECT_ID]/library/. renv.lock written. Verification passes. | [ ] Pass [ ] Fail |
| 2 | Run ./admin/admin_install_Python from project root | All Python packages installed at pinned versions in ~/.venvs/[PROJECT_ID]/. Verification passes. | [ ] Pass [ ] Fail |
| 3 | Run ./admin/admin_validate from project root | VALIDATION RESULT: PASSED. Evidence file written to ~/.jrscript/[PROJECT_ID]/validation/. R and Python versions confirmed. All packages at pinned versions. | [ ] Pass [ ] Fail |

## 4.3 IQ Evidence

Attach the evidence file generated by admin_validate as Appendix A of the Validation Report. The evidence file path is:

~/.jrscript/[PROJECT_ID]/validation/validation_[DATETIME].txt

# 5. Operational Qualification (OQ)

The OQ confirms that the JR Validated Environment functions as designed under normal and boundary conditions. Tests include destructive actions (environment deletion, file modification) that cannot be automated and require human execution and observation.

Important: The OQ must be executed after all functional development is complete. Any functional change after OQ invalidates this plan and requires a full re-run.

## 5.1 OQ Test Environment

| Hardware | Macbook Air, Apple M3, 16Gb |
| --- | --- |
| macOS version | macOS Sequoia 15.7.3 |
| R version | R 4.5.0 |
| Python version | Python 3.11.9 |
| JR env version | v1.0 (git tag: v1.0) |
| Tester | J. Rous |
| Test date | 2026-02-12 |

## 5.2 OQ Test Protocol

| ID | Description | Steps | Expected Result | Result | Findings | Notes |
| --- | --- | --- | --- | --- | --- | --- |
| OQ-01 | R environment rebuild from scratch | 1. Delete ~/.renv/[PROJECT_ID] 2. Run admin_install_R 3. Verify output | Packages installed at pinned versions in ~/.renv/[PROJECT_ID]/library/. Verification passes. | PASS | 3 findings — see Section 6 | FINDING: renv::restore() installed into system library, not isolated library. FIX: Replaced with install.packages(lib=lib_path). FINDING: renv::paths$library() resolved incorrectly. FIX: file.path(renv_root, "library"). FINDING: dir.create() missing before renv.lock write. FIX: Added. |
| OQ-02 | Python environment rebuild from scratch | 1. Delete ~/.venvs/[PROJECT_ID] 2. Run admin_install_Python 3. Verify output | Packages installed at pinned versions. Verification passes. | PASS | None | No issues found. |
| OQ-03 | Run R hello-world script | 1. Run: jrrun jrhello.R "Hello" 2. Verify output and plot | Script executes in validated R environment. Output and plot generated without error. | PASS | None |  |
| OQ-04 | Run Python hello-world script | 1. Run: jrrun jrhello.py "Hello" 2. Verify output and plot | Script executes in validated Python environment. Output and plot generated without error. | PASS | None |  |
| OQ-05 | admin_validate full IQ validation | 1. Run admin_validate 2. Verify PASSED result 3. Verify evidence file written | VALIDATION RESULT: PASSED. Evidence file generated. R and Python versions and all packages confirmed. | PASS | None after OQ-01 fixes | Passed after OQ-01 fixes were applied. |
| OQ-06 | Auto-rebuild on requirements change | 1. Delete hash file in ~/.renv/[PROJECT_ID]/ 2. Run jrrun jrhello.R 3. Verify auto-rebuild triggered | jrrun detects missing hash, triggers rebuild, then executes script successfully. | PASS | None |  |
| OQ-07 | R --add package argument | 1. Run admin_install_R --add jsonlite==2.0.0 2. Verify package added 3. Verify R_requirements.txt updated | Package added to local repo and installed. R_requirements.txt updated. Verification passes. | PASS | 1 documented limitation | LIMITATION: miniCRAN fetches current CRAN version, not pinned version. Initial attempt with ==1.8.9 installed 2.0.0. Re-run with ==2.0.0 passed. The local repository is the version control artefact for R. |
| OQ-08 | Python --add package argument | 1. Run admin_install_Python --add requests==2.31.0 2. Verify package and dependencies added 3. Verify python_requirements.txt updated | Package and all transitive dependencies installed at exact pinned versions. python_requirements.txt updated. | PASS | None | pip correctly pinned all transitive dependencies: certifi, charset-normalizer, idna, urllib3. |
| OQ-09 | Integrity check — tampered file | 1. Modify a tracked file (e.g. add whitespace to bin/jrrun) 2. Run jrrun jrhello.R 3. Restore file and re-run | Step 2: PROJECT INTEGRITY CHECK FAILED error. Step 3: runs successfully. | PASS | None |  |
| OQ-10 | Bypass protection — direct script invocation | 1. Run: Rscript R/jrhello.R "Hello" 2. Run: python3 Python/jrhello.py "Hello" | Both commands fail without executing in the validated environment. | PASS | 1 documented asymmetry | R: explicit RENV_PATHS_ROOT check with clear error referencing jrrun. Python: fails at import (no module). Asymmetry noted — Python improvement deferred to v1.1. |

# 6. Findings Summary

The following findings were raised during OQ execution. All findings were resolved prior to OQ completion. No open findings remain.

| ID | Test | Severity | Description | Resolution | Status |
| --- | --- | --- | --- | --- | --- |
| F-01 | OQ-01 | Major | renv::restore() installed packages into system library, not the isolated validated library. Environment isolation not functioning. | Replaced renv::restore() with install.packages(lib = lib_path). Added lib.loc = lib_path to verify step. | Closed |
| F-02 | OQ-01 | Major | renv::paths$library() resolved relative to working directory, not RENV_PATHS_ROOT. Validation script pointed to wrong library path. | Replaced renv::paths$library() with file.path(renv_root, "library") in admin_generate_validate_R. | Closed |
| F-03 | OQ-01 | Minor | admin_R_install.R failed to write renv.lock after clean environment deletion because target directory did not exist. | Added dir.create(dirname(lock_path), recursive=TRUE, showWarnings=FALSE) before writeLines call. | Closed |

## 6.1 Documented Limitations

| ID | Test | Description | Risk Assessment | Action |
| --- | --- | --- | --- | --- |
| L-01 | OQ-07 | miniCRAN fetches current CRAN versions during --rebuild and --add, not pinned versions from R_requirements.txt. R_requirements.txt reflects what is in the repo. | Low. The local repository is the controlled artefact. Rebuilding from CRAN requires re-validation. Normal operation uses the frozen local repo only. | Accept |
| L-02 | OQ-10 | Python scripts invoked directly outside jrrun fail at import time rather than displaying an explicit error message referencing the jrrun wrapper. | Low. Scripts cannot execute without validated packages regardless. User still cannot bypass the validated environment. | v1.1 |

# 7. Performance Qualification (PQ)

The PQ confirms that the JR Validated Environment performs correctly in realistic use — running actual analysis scripts and verifying outputs against known expected results.

## 7.1 PQ Test Protocol

| ID | Script | Procedure | Expected Result | Result | Actual Output |
| --- | --- | --- | --- | --- | --- |
| PQ-01 | jrhello.R | Run: jrrun jrhello.R "Validation" and capture output | Script runs in validated R environment. Prints message and generates plot without error. Exit code 0. | PASS | .png file showing the string “Validation” |
| PQ-02 | jrhello.py | Run: jrrun jrhello.py "Validation" and capture output | Script runs in validated Python environment. Prints message and generates plot without error. Exit code 0. | PASS | Shows plot with animation and the string “Validation” |

# 8. Acceptance Criteria

The JR Validated Environment v1.0 is considered validated when all of the following criteria are met:

| # | Criterion | Met? |
| --- | --- | --- |
| 1 | All IQ steps pass — correct R and Python versions, all packages at pinned versions | [ X] Yes [ ] No |
| 2 | All OQ tests pass or have documented accepted findings | [ X] Yes [ ] No |
| 3 | All OQ findings are closed or formally accepted as documented limitations | [ X] Yes [ ] No |
| 4 | PQ-01 and PQ-02 pass — hello-world scripts execute correctly in validated environments | [ X] Yes [ ] No |
| 5 | IQ evidence file attached to Validation Report as Appendix A | [ X] Yes [ ] No |
| 6 | Validation Report reviewed and signed off by Validation Manager | [X ] Yes [ ] No |

# 9. Document Control

## 9.1 Version History

| Version | Date | Author | Description |
| --- | --- | --- | --- |
| 1.0 | 2026-03-12 | Joep Rous | Initial release — v1.0 validation plan |

## 9.2 Related Documents

- jr_validation_report_template.docx — Validation Report template

- admin_manual.docx — Administrator Manual

- CHANGELOG.md — Software change history

- admin/project_integrity.sha256 — Project integrity file

## 9.3 Final Sign-off

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| Validation Manager | Joep Rous | JR | 2026-03-12 |